Spectroscopy using a monochromator.

You will examine several sources of light, treating some as unknowns and measure properties of the spectra produced.  There are many tools that come together to do this.

Rydberg Constant:  One task you will have is to determine the Rydberg constant for the hydrogen Balmer series spectra (see below).  Tasks will be further described below. In order to examine spectra you will need a source of light (here things like Hydrogen or Mercury discharge lamps (ask), or our sodium lamp), also a spectrometer (scanning diffraction grating monochromator), and a detector (photomultiplier tube).

The equipment involves three primary elements.

- Scanning Monochromator.  This is a diffraction grating spectrometer (1800 grooves/mm holographic grating).  Light enters one slit, diffracts off a grating which disperses the different colors of light, and depending on the tilt of the grating a single color makes it to the exit slit.  This is a computer controlled device.  You should not need to use the scanning control at all—only use the control program on the computer.  Thanks to student Jaime O’Meara, we have a newer version of the program.  You may also need to adjust entry and exit slit widths, and optical alignment. The size of the adjustable slits determines the resolution of the instrument. The monochromator should be used with the slits reduced to near their smallest setting (as long as there is sufficient signal).  A reading of zero really means 3 microns since the sharp edged slits do not close on themselves all the way (to prevent damage).  Each mark on the micrometers represents 10 microns.  So five of the smallest markings opens the slits to 50 microns (plus 3 which may be roughly ignored).  Note that such micrometer devices have offsets (non-perfect zero readings) so you may need to adjust the slits to see (observe signal) when they start opening.

When set at zero nm the instrument should send through a maximum signal since zero order reflection reaches the exit slit (like the zero order reflection on Schawlaow’s ruler).  There may be (is ) an offset on wavelength readings which you will need to account for.  You may scan through wavelengths from about negative 1.00nm to 1.00 nm to check where zero truly is.  You may want to do this manually to find the location where the peak reading is maximum (true zero) so you can account for instrument offset.  YOU WILL NEED TO ACCOUNT FOR THIS OFFSET IN YOUR DATA. When you make graphs and analyze hydrogen spectral lines, I expect that you will have already taken into account your measured offset (in Origin you can use “set column values” to do spreadsheet math).

Setting the slits at a small 50 microns gives about the highest resolution for the instrument, however very little light gets through at such settings.  You will need to think carefully about scan speeds and range to use the instrument effectively.  You may also need to use the slits to adjust for more or less signal.

A lens accompanies the setup and is selected to be roughly the correct f number for the instrument.  This means that an input angle for light matches the instrument to fill the diffraction grating.  Over filling the grating dumps light inside the instrument which scatters and gives background signal (noise).  Under filling the grating uses fewer diffraction grooves and hence gives broader diffraction maximums, hence less resolution.  I’ll leave it to you to form a good image on the slits and record/decide upon both lens and lamp positions.  A light source will need to be an appropriate distance to from an image on the slits.     You should make sure that you are able to see an image of your light source well centered on the slits. Since you will need to move light sources you will need to check on the optical alignment (object—lens—image).

Also note, that uneven settings on the slit size may be noticeable on your scans.  You may consult with other classes (Computational) regarding the effect of convolution of the slit image on the grating pattern.

DO NOT SET THE PHOTOMULTIPLIER TUBE POWER SUPPLY OVER 1200 VOLTS.  Usually less will be needed.   DO NOT PLACE MONOCHROMATOR CONTROLLER FEET DIRECTLY ON METAL TABLE SURFACE.

- A detector is mounted on the exit of the monochromator.  There is also a mechanical shutter (up is open, down closed). The photomultiplier tube (PMT) is a very sensitive device which is capable of detecting single photons and with a layer that emits electrons for each photon, and then many additional stages that emit many electrons for each electron striking that stage. The current produced is detectable with other instruments such as a lockin amplifier or piccoammeter (item 3).  The photomultiplier requires powering in order to amplify.  The setting on the power supply for the PMT is often set at -1000Volts.  In general do not adjust beyond -1200.   Lower voltages may be used. However if one is doing spectroscopy where the intensity must be known, then the voltage must remain constant throughout the experiment.  We are not doing intensity dependent spectroscopy, but rather are merely finding the wavelengths (not relative intensities) of spectral features.  The response of PMT to different colors is also an issue to consider (quantum efficiency of the detector).  You may need to check and adjust the PMT power supply voltage to increase or reduce your signal level—however you may use negative 1.00kVolt as a good benchmark, and also use monochromator slits to adjust signal.  Under no circumstances should you open the monochromator or the PMT housings.  If you notice an overload signal on the Lockin amplifier (red lights) then reduce the slit width first, and the PMT voltage second.

No oil, no food, no dirty hands, no drinks,

- The third instrument we use (other than our light source and a computer) is the SR810 digital lockin amplifier.  Think of this as a very fancy voltmeter.  In our case we will set it to read small currents.  In general the output of the PMT should not be greater than about 1.00 microAmps.  Using the lockin amplifier, a signal of a few nA is detectable.  The usable signal on the lockin will probably be on the order of tens of nanoAmps (give or take an order of magnitude). A chopper wheel is used to introduce a known reference frequency into the desired signal (the chopper must be placed in the light path between the lens and entry slits), the lockin is then able to reject noise which does not have that same frequency (thermal, background light, electrical, etc).  The lockin is a very  powerful piece of equipment.  The chopper may be set at roughly a few hundred Hz (as long as not a multiple of 60Hz line noise)—lets say 280Hz.  The phase between the chopper reference must be synchronized with the signal—hit “autophase” once you have some signal.  The source should be on current I x106 and you will need to select the appropriate sensitivity scale (this depends on your signal level).  This is like setting the number of millivolts/box on an oscilloscope.  The time constant (think RC smoothing and also FFT integration time) is adjustable ---as a guess something on the order of 100ms should be appropriate.  Note this goes into thinking about your scan speed!!!!!!!!!!!!!  You must take at least a time constant to scan through a wavelength range which you want to resolve.  Oversampling is good.

For example:  If I scan through 0.1nm/second, but have a time constant of 20seconds, then I cannot resolve (see) features at the approx. 0.05nm resolution of the instrument.  I fact here my best resolution would be about 2nm---The sample rate, scan speed, and time constant must be thought out ahead of time and must match up.

If your monochromator can observe approx 0.05 nm, then you should not scan faster than 0.05nm in 0.1 s (the time const) if you want full resolution.  So in general with this example (100 ms time constant), never scan faster than 30nm/minute on the monochromator.   There are many other settings on the lockin apmplifier which will be preset to examine signal in these experiments.

As you take data you will end up with scans of intensity (proportional to intensity) vs wavelength.  These will be sent as data with both intensity and wavelength in a *.csv  or *.txt file---which origin or other software can import.  There will be big picture scans (fast through entire spectrum) and also zoom in scans on a narrow feature  (slow through individual spectral lines).   Your goal is to determine the wavelength of various spectral features (peaks).  The files are saved to the desktop and they can be opened and viewed with origin right click open with…..   Copy to your flash drives.

The goal of fast scans is to see where specific intense spectral features are located.  So for example, in hydrogen you will want to scan from about 400nm to 700nm and look for 4 intense features.  Then you will want to scan slowly through each feature and get a high resolution (slow with lots of samples) scan so that you can determine the peak wavelengths.

For each spectral feature you analyze you will make a graph which you can analyze to determine a peak wavelength and a linewidth.   This width is a convolution of the instrument resolution and the true width of the spectral feature.  For example –the width of a HeNe laser line (light scattered off a note card into the monochromator) should appear to be a delta function with very narrow width, but will appear to be about 0.05 to 0.1 nm due to the limitations of the instrument resolution.  Several factors influence the width and the intensity of spectral features.

For each of the graphs below (except graph 1), account for any instrumental offset so that the wavelength reading is your best true number.

TASKS

- With any source determine the zero wavelength offset for the monochromator. (find the peak near zero   it should be within a few tenths of nm).  You will have this graph in analysis and final draft.  GRAPH 1.  Ideally there should be no geometric changes to optics after this, but check.

- Verify that the sodium lines are where you expect (given any offset).  Both these lines will fit on a single scan.  You know the numbers from previous lab.  Measure the wavelength separation and compare to previous lab.   GRAPH 2

- Detailed long range and narrow focus scans for the 4 visible Balmer series spectral lines (long slow scan should tell you where to look, then zoom in for each spectral line).   You will need to have a single long range scan (plot) that has all four Balmer series lines.   Also four individual scans (plots) with individual lines.  For each individual line you need to measure the peak wavelength (including any offset) and the width of the peak.  GRAPHS 3,4,5,6,7 (ONE LONG AND 4 LINES)

- Your analysis will be to make an appropriate plot to determine the Rydberg Constant. (another plot).  GRAPH 8.

- Other—above and beyond ideas----Things for you to consider ? Other source (ask)?  Instrument resolution (measure and change)?  Detailed discussion of analysis and spectral features?   relative intensity and instrument intensity calibration?  (How tall should each peak be?)   Absorption spectra—of a sample, ???  ?  Properties of diffraction gratings in such spectrometers.  Instrument control.   How does lockin amplifier work.   Deconvolution to determine instrument profile?  Etc.

You will use the scanning monochromator, a photomultiplier tube (very sensitive detector), and a lockin amplifier.  You will examine the spectral output from a Hydrogen discharge tube.  The tube emits light due to excitation of hydrogen atoms and also hydrogen molecules.  The geometry of the tube is set up so that the electrical discharge eliminates most of the H2, thus reducing the relative intensity of the molecular spectra.  What remains is largely hydrogen Balmer series line emission.

You will scan and observe the peak wavelengths for the Balmer series lines in Hydrogen.  Observe as many as possible (four are likely).  The Balmer series has transitions from upper states (n) to a final state of 2.   (Eq 11.2 page 126 , Essentials of Modeer Physics, T.R. Sandin).  You can look up in any reference you wish.

The constant displayed is the Rydberg constant.  You will use your data to determine this constant.  This requires measurement of the wavelengths for several peaks of intensity output from the hydrogen lamp.  Then you will make an appropriate plot to determine the Rydberg Constant.

Remember to use the sensitivity adjustment on the lockin, or the monochromator slits, or the PMT voltage to ensure that the signal does not overload the lockin amplifier.